| Initiated by: Masab Rasheed | Department: QC Lab | CAR No. CA-03-02-2025 |
| --- | --- | --- |
| Designation: AQCM | CAR Initiated Date: 03-02-2025 | NC No. 03 |

CAR Initiated due to (Please tick the relevant appropriate box below)

PNAC Assessment                ☐ Complaint                 ☐ Accident/Incident              ☐  Internal Audit          ☐ Suggestion/Improvement      ☐  Technical Fault        ☐  Others

| - Description (to be filled by Initiator) |  |  |
| --- | --- | --- |
| All the equipment/apparatus have unique identification number and calibration sticker, wherever required and the equipment logbook analytical balance (QC-INS-001) not maintained (NC). The Calibration plan (QCL-FRM-4.04) for the year 2024 is available in file no (QCL-FIL-22). the lab has developed the Equipment maintenance records sheet (QCL-FRM-4.03) and records related to Equipment maintenance is not maintained (NC). ____________ Signatures |  |  |
| Significance/Priority Level | Urgent ☐ Normal ☐ Non-significance |  |
| Accepted ☐ Rejected | Assigned to: Masab Rasheed Date: 03-02-2025 |  |
| * If accepted then Manager QC to mark it to concerned person for Root Cause Analysis. |  |  |
| - Root Cause Analysis (where required) |  |  |
| - The lack of a systematic process for ensuring logbook entries is due to insufficient awareness and training on LSP adherence, coupled with inadequate oversight and accountability mechanisms within the Laboratory Management System (LMS). - The absence of a systematic review and training program on the maintenance procedure and record-keeping requirements. |  |  |
| - Proposed Action | Target Date: | 07-02-2025 |
| Corrective action for shortcomings on Both NCs has been taken by maintaining the log book for analytical weighing balance (QCL-LB-07). Lab has started to maintain the equipment maintenance records. _____________________ Name or Signatures (Person who carried out Root Cause Analysis and Proposed Corrective Action) |  |  |
| - Verification by Manager QC for implementation & Effectiveness |  |  |
| Is Proposed Action Implemented? YES ☐ NO |  | Is it Effective? YES ☐ NO |
| Remarks (If any): CAR closing date: __________________________ Manager QC Sign: _________________________ |  |  |