| Initiated by: Masab Rasheed | Department: QC Lab | CAR No. CA-04-02-2025 |
| --- | --- | --- |
| Designation: AQCM | CAR Initiated Date: 03-02-2025 | NC No. 04 |

CAR Initiated due to (Please tick the relevant appropriate box below)

PNAC Assessment                ☐ Complaint                 ☐ Accident/Incident              ☐  Internal Audit          ☐ Suggestion/Improvement      ☐  Technical Fault        ☐  Others

| - Description (to be filled by Initiator) |  |  |
| --- | --- | --- |
| The Lab has developed a procedure related to metrological traceability (QSP-LSP-05) The reference standards used are valid and traceable to their primary standards and stored in refrigerator but monitoring the temperature through un calibrated thermo hygrometer (NC). ____________ Signatures |  |  |
| Significance/Priority Level | Urgent ☐ Normal ☐ Non-significance |  |
| Accepted ☐ Rejected | Assigned to: Masab Rasheed Date: 03-02-2025 |  |
| * If accepted then Manager QC to mark it to concerned person for Root Cause Analysis. |  |  |
| - Root Cause Analysis (where required) |  |  |
| The procedure (QSP-LSP-05) does not include the requirement for calibration of auxiliary monitoring equipment, such as the thermo-hygrometer, leading to its use without calibration. |  |  |
| - Proposed Action | Target Date: | 18-02-2025 |
| Corrective action for shortcomings on NC has been taken by the calibration of thermo hygrometer for monitoring the temperature. _____________________ Name or Signatures (Person who carried out Root Cause Analysis and Proposed Corrective Action) |  |  |
| - Verification by Manager QC for implementation & Effectiveness |  |  |
| Is Proposed Action Implemented? YES ☐ NO |  | Is it Effective? YES ☐ NO |
| Remarks (If any): CAR closing date: __________________________ Manager QC Sign: _________________________ |  |  |