| Non-Conformance Description |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| The Lab has developed a procedure related to metrological traceability (QSP-LSP-05) The reference standards used are valid and traceable to their primary standards and stored in refrigerator but monitoring the temperature through un calibrated thermo hygrometer (NC). |  |  |  |  |  |  |  |  |
| RCA Team Members |  |  | Himayat Hussain (QCM), Masab Rasheed (AQCM) |  |  |  |  |  |
| RCA Team Lead / Head |  |  | Himayat Hussain (QCM) |  |  |  |  |  |
| RCA Start Date |  |  | 03-02-2025 |  | Submission Date |  | 18-02-2025 |  |
| Methodology / Tools Used |  |  |  |  |  |  |  |  |
| Cause & Effect Diagram | □ | 5WHY analysis |  | □ |  | 3W2H / 5W2H |  |  |

Cause Effect Diagram

Documentation

Training

Root Cause

Proper tools

Method

| 3W2H /5W2H | Phenomenon of Problem / Question | Explanation / Answer |
| --- | --- | --- |
| WHY | Why was the temperature monitoring of the refrigerator using an uncalibrated thermo-hygrometer? | There was no calibration schedule or tracking mechanism for the thermo-hygrometer. |
| WHY | Why was there no calibration schedule for the thermo-hygrometer? | Calibration of auxiliary monitoring devices (like thermo-hygrometers) was not considered a priority or overlooked during procedure development. |
| WHY | Why was calibration of the thermo-hygrometer overlooked? | The focus was on ensuring traceability and validity of primary reference standards, but monitoring tools (support equipment) were not included in the calibration scope. |
| WHY | Why was monitoring tools not included in the calibration scope? | There was a gap in the procedure (QSP-LSP-05) that did not specify the need for calibration of monitoring equipment like thermo-hygrometers. |
| WHY | Why was there a gap in the procedure (QSP-LSP-05)? | The risk assessment during the procedure's development did not account for the potential impact of uncalibrated monitoring equipment on traceability. |
| How | How to address this specific NC (Corrective Action)? | Identify all monitoring equipment (like the thermo-hygrometer) used in the lab and send them for calibration to an accredited body. Update the refrigerator monitoring records with data from the calibrated equipment. |
| How | How to prevent this NC from recurring (Preventive Action)? | - Revise the procedure (QSP-LSP-05) to include requirements for periodic calibration of all monitoring equipment. - Conduct training sessions for relevant staff on the importance of calibrating auxiliary equipment. - Implement a system to track calibration schedules for all lab equipment, including monitoring devices. |
| Root Cause: The procedure (QSP-LSP-05) does not include the requirement for calibration of auxiliary monitoring equipment, such as the thermo-hygrometer, leading to its use without calibration. |  |  |

RCA Statement: The procedure (QSP-LSP-05) does not include the requirement for calibration of auxiliary monitoring equipment, such as the thermo-hygrometer, leading to its use without calibration.

Reviewed by: _______________________	                           Evaluated by: ______ ______