| NC# | Date: | Non-Conformance | Source | Initiated By |
| --- | --- | --- | --- | --- |
| 145 | 23-01-2023 | The Lab has not addressed the requirement of the standard regarding taking practical measures to protect unintended adjustment to equipment , as stipulated vide clause-6.4.12 of the standard(NC-01) | Ex. Audit | M. Asif Rana |
| 146 | 23-01-2023 | The Lab has not re-evaluated the external providers of products & services as required by clause -6.6.2(d) of the standard ISO/IEC 17025:2027(NC-02a). | Ex. Audit | M. Asif Rana |
| 147 | 23-01-2023 | During the last assessment ,it was observed and advised by the TL , to include and use all possible internal quality assurance tools for ensuring the validity of results , as required by the standard ,but the Lab has not taken appropriate action to address the observation(NC-03). | Ex. Audit | M. Asif Rana |
| 148 | 23-01-2023 | The Lab has no documented record relating to acknowledgement of policies and objectives, at all levels of Lab organization, for implementation, as required by clause- 8.2.1 of the standard ISO/IEC 17025:2017(NC-04a). ii)During the last assessment the Lab was advised to identify in the Doc review record, the personnel who reviewed the documents, but the Lab has not taken any CA in this regard(NC-04b). The record of change made to a document #PCC-PTL-LSP-014 , Rev.01,titled as “LSP for Sampling /Handling of Test Items vide DCR #2022-37,dated 22-10-2022, was checked and observed the Lab has not identified the changed text in the revised document as described in Lab’s own procedure (NC-04c) | Ex. Audit | M. Asif Rana |
| 149 | 23-01-2023 | During the last held internal audit, the auditors did not use the internal audit check list developed by the Lab vide Doc# PCC-PTL/LSF-8.03, as described in Lab’s own procedure procedure(NC-05). | Ex. Audit | M. Asif Rana |
| 150 | 23-01-2023 | The Lab has not addressed in the output minutes of the meeting ,the requirements of the standard as mentioned at clause-8.9.3 of the standard ISO/IEC 17025 :2017 (NC-06) | Ex. Audit | M. Asif Rana |
| 151 | 23-01-2023 | Although Lab has update procedure for defining criteria of Selection of Personnel and Authorization of Personnel. However, Lab has not developed procedure for selection of the trainer (NC). the lab has developed a new inducted staff training farm (PCC-PTL-LSF-10.10) But record related to the training is not available(NC). Lab has also not defined the training evaluation criteria of the new staff (NC). Although Lab has developed Annual Training Program for year 2022 (PCC-PTL-LSF-10.03). But not define Internal training need of lab Staff (NC). | Ex. Audit | M. Asif Rana |
| 152 | 23-01-2023 | The lab has also one sample retain room and monitoring the temperature and humidity by using thermo hygrometer (PCC/THG/06).(NC). The lab has not authorized the person for the monitoring the temperature and humidity (NC). | Ex. Audit | M. Asif Rana |
| 153 | 27-05-2023 | GC-7890B system is not communicating with software system. | Lab Activity | M. Asif Rana |
| 154 | 14-10-2023 | The Lab has addressed in the procedure at clause 4.4, about the periodic review of LMS documents to ensure the continuing suitability with the applicable requirements, of the standard ISO/IEC 17025:2017. For this purpose the Lab has develop “Document Periodic Review Plan ‘ vide Doc# PCC-PTL-LSF-1.08 and record of “Annual Review of Documents : vide Doc# PCC-PTL-LSF-1.07. The work instructions and external origin documents are missing in the Document Periodic review plan. (Hence NC is raised) | Int. Audit | M. Asif Rana |
| 155 | 14-10-2023 | Since the last assessment, the Lab has taken thirty-two (32) corrective actions (CAs). All the CAs have been closed except the last one which is under process, as checked for compliance and observed that the Lab has carried out RCA using description of four indicators to analyze the root cause following the fish bone statistical tool. But diagram of statistical tools (Fish Bone) is missing in procedure. (Hence NC is raised) | Int. Audit | M. Asif Rana |
| 156 | 17-10-2023 | The PCC-PTL has documented the temperature and humidity requirements of instrument rooms, sample preparation & physical analysis room (20-25⁰C and 35 - 65%rh) in “LSP for Facilities and Environmental Conditions” (Doc # PCC-PTL-LSP-011). For Refrigerator 15±100C Temperature and Humidity 40±15 is defined. But Lab has not defined Environmental Condition for Sample Storage room in procedure “LSP for Facilities and Environmental Conditions” (Doc # PCC-PTL-LSP-011). (Hence NC is raised | Int. Audit | M. Asif Rana |
| 157 | 17-10-2023 | Equipment maintenance is followed by system performance check as mention in SOP for equipment management LSP-013, One record for auto sampler maintenance is checked but system performance was not carried out after maintenance. (Hence NC is raised) | Int. Audit | M. Asif Rana |
| 158 | 17-10-2023 | The lab is making Decision in terms of Fit for use or unfit for use. The lab has developed the decision rule for reporting result (PCC-PTL-LSF-2.04). The lab has not documented statement of Conformity considering the level of risk associated with decision rule as per ILAC G-08 applied in their results. (NC). Lab need to mention the decision rule in QM or in procedure. (Hence NC is raised) | Int. Audit | M. Asif Rana |
| 159 | 06-12-2023 | We received complaint from our SukKhar Zone Regarding solubility of Timer Gold 3% SC (Batch # PC23007ELSN, Mfg: Nov, 2023 Exp: Nov, 2025) as product has dissolution issue. We what’s up the pictures and video as an evidence shared by the customer. Kindly investigate the solubility issue and share purpose solution. | Customer complaint | M. Asif Rana |
| 160 | 08-12-2023 | We received complaint from our Sahiwal Zone regarding SUCCESS 72 % WP (Batch # PC21009RSCPO, Mfg: Nov, 2023 Exp: Nov, 2025) as product has burning issue/ phytotoxic effects on plant. We shared the pictures as an evidence received from the customer. Kindly reanalyze batch retained sample to investigate the burning issue/Phytotoxic effects. | Customer complaint | M. Asif Rana |
| 161 | 07-02-2024 | Liquid Chromatogram (PCC-PTL-HPLC-02) pump is not getting on. There may be a power supply issue. | Lab Activity | M. Asif Rana |

Prepared By:________________								updated on:______________________