Date: 11-02-2025					             NC #: 10

| - Non-Conformance Detail |  |  |
| --- | --- | --- |
| 01 | Non-Conformance Description | Laboratory has not applying the decision rule and also not mentioned lab procedure of reporting of results (NC). Lab has not applying the decision rule for the statement of conformity (NC). The VAN laboratory is providing interpretation like ok but the criteria regarding ok and not ok not mentioned in procedure (QCL-LSP-19) (NC). The lab has a procedure for reporting of results but not mentioned about amended/duplicate report issuance not mentioned in procedure (NC). |
| 02 | Non-Conformance Source | Audit NC □ Laboratory Activities □ Suggestion □ PT/ILC □ Complaint □ Customer Feedback □ Risk Assessment □ Any Other___________________ |
| 03 | Non-Conformance Initiated By | Masab Rasheed |

| - Non-Conformance Significance Evaluation |  |
| --- | --- |
| Impact on Previous Result | □ Yes No |
| Level of Risk | □ Very Low □ Low Moderate □ High □ Very High |
| Non-Conformance Type | Essential □ Minor Non- Essential |
| Acceptance of Non-Conformance | Accepted □ Rejected |

| - Actions required on Non-Conformance |  |
| --- | --- |
| Withhold of reports | □ Yes No |
| Halt of work | □ Yes No |
| Recall of work | □ Yes No |
| Revision of Documents | Yes □ No |

| Reviewed By: Masab Rasheed AQCM | Evaluated by: Himayat Hussain QCM |
| --- | --- |
| Signature: | Signature: |
| Date: | Date: |