| Initiated by: Masab Rasheed | Department: QC Lab | CAR No. CA-02-02-2025 |
| --- | --- | --- |
| Designation: AQCM | CAR Initiated Date: 03-02-2025 | NC No. 02 |

CAR Initiated due to (Please tick the relevant appropriate box below)

PNAC Assessment                ☐ Complaint                 ☐ Accident/Incident              ☐  Internal Audit          ☐ Suggestion/Improvement      ☐  Technical Fault        ☐  Others

| - Description (to be filled by Initiator) |  |  |
| --- | --- | --- |
| The lab is monitoring the temperature and humidity by using thermo hygrometer (VAN/PFTL/EQP-001) in main lab hall while in sample storage room by using the Thermo hygrometer (VAN/PFTL/EQP-002) all are calibrated but the lab has not authorized the person for this activity (NC). The lab has not developed and maintained the log books for analytical weighing balance (QCL-INS-01) and Block heater (QC-INS-023) (NC). ____________ Signatures |  |  |
| Significance/Priority Level | Urgent ☐ Normal ☐ Non-significance |  |
| Accepted ☐ Rejected | Assigned to: Masab Rasheed Date: 03-02-2025 |  |
| * If accepted then Manager QC to mark it to concerned person for Root Cause Analysis. |  |  |
| - Root Cause Analysis (where required) |  |  |
| - Lack of a formal process or system for documenting and authorizing personnel responsible for monitoring equipment, resulting from insufficient awareness or oversight regarding the quality management system requirements for personnel authorization. - Lack of awareness and clarity among responsible personnel regarding the requirement to develop and maintain logbooks for analytical weighing balance (QCL-INS-01) and block heater (QC-INS-023), primarily due to gaps in training, LSP communication, and oversight mechanisms. |  |  |
| - Proposed Action | Target Date: | 08-02-2025 |
| Corrective action for shortcomings on Both NCs has been taken by revising the Authorization permit of Lab Assistant (QCL-FRM-2.01). Lab has maintained the log books for analytical weighing balance (QCL-LB-07) and Block heater (QCL-LB-08). _____________________ Name or Signatures (Person who carried out Root Cause Analysis and Proposed Corrective Action) |  |  |
| - Verification by Manager QC for implementation & Effectiveness |  |  |
| Is Proposed Action Implemented? YES ☐ NO |  | Is it Effective? YES ☐ NO |
| Remarks (If any): CAR closing date: __________________________ Manager QC Sign: _________________________ |  |  |