| Initiated by: Masab Rasheed | Department: QC Lab | CAR No. CA-01-05-2026 |
| --- | --- | --- |
| Designation: AQCM | CAR Initiated Date: 11-05-2026 | NC No. 01 |

CAR Initiated due to (Please tick the relevant appropriate box below)

☐ PNAC Assessment                ☐ Complaint                  Accident/Incident              ☐  Internal Audit          ☐ Suggestion/Improvement      ☐  Technical Fault        ☐  Others

| - Description (to be filled by Initiator) |  |  |
| --- | --- | --- |
| Blast of KJeldhal’s Assembly Splash head during Nitrogen Testing. ____________ Signatures |  |  |
| Significance/Priority Level | Urgent ☐ Normal ☐ Non-significance |  |
| Accepted ☐ Rejected | Assigned to: Masab Rasheed Date: 11-05-2026 |  |
| * If accepted then Manager QC to mark it to concerned person for Root Cause Analysis. |  |  |
| - Root Cause Analysis (where required) |  |  |
| Lack of standardized verification check list of Kjeldhal’s assembly setup before operation. |  |  |
| - Proposed Action | Target Date: | 12-05-2026 |
| Replace damaged splash heads/glassware immediately and the equipment was functionally verified before further use. Prepared pre-use inspection check list and laboratory personnel were instructed on proper equipment handling and operational checks. Additionally, check list display on board for routine verification before testing nitrogen on Kjeldhal’s Apparatus. _____________________ Name or Signatures (Person who carried out Root Cause Analysis and Proposed Corrective Action) |  |  |
| - Verification by Manager QC for implementation & Effectiveness |  |  |
| Is Proposed Action Implemented? YES ☐ NO |  | Is it Effective? YES ☐ NO |
| Remarks (If any): CAR closing date: __________________________ Manager QC Sign: _________________________ |  |  |