| Initiated by: Masab Rasheed | Department: QC Lab | CAR No. CA-06-07-2026 |
| --- | --- | --- |
| Designation: AQCM | CAR Initiated Date: 22-07-2026 | NC No. 03 TA |

CAR Initiated due to (Please tick the relevant appropriate box below)

PNAC Assessment                ☐ Complaint                 ☐ Accident/Incident              ☐  Internal Audit          ☐ Suggestion/Improvement      ☐  Technical Fault        ☐  Others

| - Description (to be filled by Initiator) |  |  |
| --- | --- | --- |
| The lab is using refrigerator for keeping the CRMs and working standards and has also maintained the records of temperature range in between (-5C to +5C) for the but the thermo hygrometer range of calibration/ sensitivity of the thermohygrometer is not intact (NC). ____________ Signatures |  |  |
| Significance/Priority Level | Urgent ☐ Normal ☐ Non-significance |  |
| Accepted ☐ Rejected | Assigned to: Masab Rasheed Date: 22-07-2026 |  |
| * If accepted then Manager QC to mark it to concerned person for Root Cause Analysis. |  |  |
| - Root Cause Analysis (where required) |  |  |
| The thermo-hygrometer was available and calibrated; however, its calibration range started above 15°C and therefore did not cover the refrigerator operating range of -5°C to +5°C. The suitability of the calibration range for the intended use was not verified before using the instrument for refrigerator temperature monitoring. |  |  |
| - Proposed Action | Target Date: | 22-08-2026 |
| The available thermo-hygrometer could not be calibrated at the required lower temperature range. Therefore, a new data logger will be purchased and calibrated to cover the refrigerator operating range of -5°C to +5°C before use for CRM and working-standard storage monitoring. _____________________ Name or Signatures (Person who carried out Root Cause Analysis and Proposed Corrective Action) |  |  |
| - Verification by Manager QC for implementation & Effectiveness |  |  |
| Is Proposed Action Implemented? YES ☐ NO |  | Is it Effective? YES ☐ NO |
| Remarks (If any): CAR closing date: __________________________ Manager QC Sign: _________________________ |  |  |