| Sr. No | Non-Conformance | Source | Root Cause | Corrective Action | Target Date |
| --- | --- | --- | --- | --- | --- |
| 01 | The Lab Assistant has not the requisite education, training and competency to conduct the test under scope which is volition of the clause 6.2.2 of the standard ISO/IEC 17025:2017 (NC). Mr. Ahmad Umair (CEO) is supervising the analytical activities in VAN lab as a whole without authorization permit (NC). The lab has developed an annual training Program me (QCL-FRM-2.03) checked the training record of Mr. Masab Rasheed authorized for training of the staff as mentioned in the Authorization permit (QCL/FRM-2.01) but he has no training over MU but working as a trainer over MU (NC). Lab has not defined the criteria of the selection of the trainer (NC). The lab has also developed a training need assessment form (QCL/FRM-2.02) but the criteria of the need assessment not mentioned in procedure as well as in training need assessment farm (QCL-FRM-2.02) (NC). The training record related to Measurement of uncertainty is not available in folder (QCL-FIL-20) (NC) however the records related to trainer evaluation is also missing (NC). | Ex. Audit | - Lack of a formal process to evaluate and verify the competency, education, and training of personnel before assigning them tasks related to specific tests under the scope of ISO/IEC 17025:2017. - Lack of a robust compliance framework and clearly defined procedures for supervisory roles, including a failure to prioritize and implement internal audits. - A lack of a robust verification mechanism for ensuring that training records (QCL-FRM-2.03) are checked against the authorization permit (QCL/FRM-2.01) before assigning staff as trainers. - The lab management failed to establish and prioritize a documented process for trainer selection due to a lack of focus on human resource competency frameworks within the Lab Management System (LMS). - The development process for the training need assessment procedure and form lacked a systematic approach to ensure that criteria for assessing training needs were clearly defined and documented. This resulted from inadequate consultation, review, and validation mechanisms during the document design and approval phases. - Lack of a standardized process or procedure for the proper documentation and filing of training records related to 'Measurement of Uncertainty' in the designated folder (QCL-FIL-20). - Lack of a defined process and accountability for documenting and storing trainer evaluation records, coupled with inadequate awareness of the importance of compliance and quality requirements. | Corrective action for shortcomings on all the NCs has been taken by revising the Authorization permit of Lab Assistant (QCL-FRM-2.01). Revised Authorization permit of CEO (QCL-FRM-2.01). Lab has conducted internal training of Masab Rasheed on Measurement Uncertainty Certificate of training attached. Lab has revised LSP (QCL-LSP-02) and defined the criteria for the selection of trainer accordingly and revised the training need assessment Form (QCL-FRM-2.02). Lab has conducted internal training on MU and maintained record related to MU in folder (QCL-FIL-20) and record related to trainer evaluation maintained in folder (QCL-FRM-20). | 13-02-2025 |
| 02 | The lab is monitoring the temperature and humidity by using thermo hygrometer (VAN/PFTL/EQP-001) in main lab hall while in sample storage room by using the Thermo hygrometer (VAN/PFTL/EQP-002) all are calibrated but the lab has not authorized the person for this activity (NC). The lab has not developed and maintained the log books for analytical weighing balance (QCL-INS-01) and Block heater (QC-INS-023) (NC). | Ex. Audit | - Lack of a formal process or system for documenting and authorizing personnel responsible for monitoring equipment, resulting from insufficient awareness or oversight regarding the quality management system requirements for personnel authorization. - Lack of awareness and clarity among responsible personnel regarding the requirement to develop and maintain logbooks for analytical weighing balance (QCL-INS-01) and block heater (QC-INS-023), primarily due to gaps in training, LSP communication, and oversight mechanisms. | Corrective action for shortcomings on Both NCs has been taken by revising the Authorization permit of Lab Assistant (QCL-FRM-2.01). Lab has maintained the log books for analytical weighing balance (QCL-LB-07) and Block heater (QCL-LB-08). | 08-02-2025 |
| 03 | All the equipment/apparatus have unique identification number and calibration sticker, wherever required and the equipment logbook analytical balance (QC-INS-001) not maintained (NC). The Calibration plan (QCL-FRM-4.04) for the year 2024 is available in file no (QCL-FIL-22). the lab has developed the Equipment maintenance records sheet (QCL-FRM-4.03) and records related to Equipment maintenance is not maintained (NC). | Ex. Audit | - The lack of a systematic process for ensuring logbook entries is due to insufficient awareness and training on LSP adherence, coupled with inadequate oversight and accountability mechanisms within the Laboratory Management System (LMS). - The absence of a systematic review and training program on the maintenance procedure and record-keeping requirements. | Corrective action for shortcomings on Both NCs has been taken by maintaining the log book for analytical weighing balance (QCL-LB-07). Lab has started to maintain the equipment maintenance records. | 07-02-2025 |
| 04 | The Lab has developed a procedure related to metrological traceability (QSP-LSP-05) The reference standards used are valid and traceable to their primary standards and stored in refrigerator but monitoring the temperature through un calibrated thermo hygrometer (NC). | Ex. Audit | The procedure (QSP-LSP-05) does not include the requirement for calibration of auxiliary monitoring equipment, such as the thermo-hygrometer, leading to its use without calibration. | Corrective action for shortcomings on NC has been taken by the calibration of thermo hygrometer for monitoring the temperature. | 18-02-2025 |
| 05 | The lab has developed the simple receiving form (QCL-FRM-12.01) and mentioned the criteria for sample acceptance/rejection like improper sample receiving, costumer/sample information is not appropriate, capability and resources do not meet the requirement, out of the detection limit of the equipment etc. but the acceptance criteria is not maintained in the procedure, lab needs to revise the procedure (QCL-LSP-22) (NC). | Ex. Audit | The lab does not have a structured review and training process for ensuring that procedural updates align with new forms and criteria. | Corrective action for shortcomings on NC has been taken by revising the procedure (QCL-LSP-22) and maintained the acceptance criteria in the procedure. | 10-02-2025 |
| 06 | The lab has not developed a procedure for validation and verification of methods (NC). The raw data regarding the method validation is not maintained (NC). | Ex. Audit | - The lab management failed to establish clear responsibilities and oversight for the development of the procedure due to insufficient awareness and training on the importance of method validation and verification as part of the quality management system. - Insufficient training and a lack of clear procedures for maintaining raw data due to inadequate management focus on compliance requirements. | Corrective action for shortcomings on both the NCs has been taken by developing the procedure for Validation and Verification of methods and maintained the raw data regarding method validation. | 14-02-2025 |
| 07 | The VAN Laboratory has not developed a lab procedure for control of technical record (QCL-LSP-20) (NC). | Ex. Audit | The laboratory lacks a structured process to review and ensure compliance with ISO/IEC 17025:2017 documentation requirements, resulting in the omission of QCL-LSP-20 for control of technical records. | Corrective action for shortcomings on the NC has been taken by developing the procedure for control of technical record (QCL-LSP-20). | 14-02-2025 |
| 08 | The lab analyst involved in reporting of Uncertainty of measurements have not sufficient knowledge over MU calculation lab staff involved in the calculation of MU needs training (NC). The raw data related to uncertainty measurements Nitrogen Phosphorus is not properly available (NC) | Ex. Audit | - The lack of a structured competency assessment and training program for MU calculations. This was due to assumptions about staff knowledge and a failure to prioritize skill development for this critical task. - Lack of a structured system to ensure procedure compliance, training, and monitoring for quality control. | Corrective action for shortcomings on both the NCs has been taken by conducting internal training of Lab Staff on Measurement of Uncertainty Certificate of training attached and raw data related to uncertainty measurements of Nitrogen, phosphorus is maintained. | 15-02-2025 |
| 09 | The lab has developed the procedure for reporting of results (QCL-LSP-19) however lab needs to revise the procedure as per Clause 7.8.1 & 7.8.2 of the standard ISO/IEC 17025:2017 (NC). | Ex. Audit | The root cause of the non-conformance is the lack of a formal process to update staff on changes in relevant standards and regulations. | Corrective action for shortcomings on the NC has been taken. Lab has revised the procedure for reporting of results (QCL-LSP-19) as per clause 7.8.1 and 7.8.2 of the standard ISO/IEC 17025:2017. | 15-02-2025 |
| 10 | Laboratory has not applying the decision rule and also not mentioned lab procedure of reporting of results (NC). Lab has not applying the decision rule for the statement of conformity (NC). The VAN laboratory is providing interpretation like ok but the criteria regarding ok and not ok not mentioned in procedure (QCL-LSP-19) (NC). The lab has a procedure for reporting of results but not mentioned about amended/duplicate report issuance not mentioned in procedure (NC). | Ex. Audit | - The laboratory staff is not consistently applying the decision rule or mentioning lab procedures for reporting results because of insufficient training, unclear communication, or lack of a formal review process to ensure adherence to these procedures. - The root cause is likely insufficient training and awareness among lab staff about the importance of applying the decision rule for the statement of conformity, compounded by inadequate communication and system checks to reinforce the correct application of the procedure. - The root cause of this non-compliance (NC) is that the procedure QCL-LSP-19 does not clearly define the criteria for labeling results as "OK" or "Not OK." Without these clear guidelines, lab staff may interpret results in different ways, leading to inconsistency and confusion. - The root cause of this non-conformance (NC) is that the lab's procedure did not include instructions on how to handle amended or duplicate reports. This happened because the procedure focused mainly on the initial report creation and didn't account for situations where reports need to be changed or duplicated. | Corrective action for shortcomings on all the NCs has been taken. Lab has revised the lab procedure for reporting of results (QCL-LSP-19) and applying the decision rule for statement of conformity. Lab has revised the procedure (QCL-LSP-19) and mentioned the criteria regarding OK and Not OK. Lab has revised the procedure for reporting of results (QCL-LSP-19) and mentioned about amended/duplicate report issuance in the procedure. | 18-02-2025 |
| 11 | The lab has also developed a format for inspection of purchased products, but has not evaluated as per standard requirements, (NC). Lab has also not defined about the re-evaluation and lab not practicing the reevaluation of vendors (NC). | Ex. Audit | - The team didn’t properly check if the inspection format met the standard requirements. This happened because there wasn’t a clear process, proper training, or oversight in place to ensure compliance. - The lab management failed to recognize the importance of re-evaluating vendors and did not include it in their processes. This happened because there was no proper training or focus on vendor management requirements during procedure development. | Corrective action for shortcomings on both the NCs has been taken by revising the format for inspection of purchased products as per standard requirements. Lab has revised the procedure (QCL-LSP-06) and defined the criteria about the re-evaluation and start practicing the reevaluation of vendors. | 07-02-2025 |
| 12 | The lab has a developed a procedure (QCL/LSP/14) for the non- conforming work. The procedure normally describes about the non-conformities. (NC) The lab has not defined responsibilities and authorities for the management of non-conforming work, (NC). The lab has developed a register for recording the non-conforming work, but lab not checking the significance. (NC) | Ex. Audit | - The root cause of this non-conformance is that the procedure (QCL-LSP-14) was designed with a narrow focus on identifying non-conformities but missed addressing how to prevent or correct them, likely due to a lack of training or awareness about comprehensive quality management practices. - The lab's management overlooked the need to define clear roles and responsibilities for handling non-conforming work, likely due to a lack of awareness or focus on this aspect during system implementation. - The root cause is incomplete training needs analysis, which led to insufficient awareness of the importance of assessing the significance of non-conforming work. | Corrective action for shortcomings on all the NCs has been taken by revising the procedure (QCL-LSP-14) and defined the responsibilities and authorities for the management of non-conforming work. The lab has revised the format on register and added the significance evaluation. | 10-02-2025 |
| 13 | The Lab has defined its quality objectives (QCL-FRM-18.01) for the year 2024. 04 quality objectives defined by the Lab with target dates & responsibilities, but the target dates of all objectives are fixed Dec 2024 without any justification and detailed roadmap. (NC) | Ex. Audit | The root cause of the nonconformance is a lack of a structured and systematic approach within the Lab's quality management system to plan, justify, and document realistic target dates with clear roadmaps for achieving quality objectives. | Corrective action for shortcomings on the NC has been taken by defining its quality objective (QCL-FRM-18.01) for year 2025 with target dates & responsibilities with detailed roadmap. | 12-02-2025 |
| 14 | The lab has developed a master list (QCL/FRM/20.02) for internal, but lab not controlled the external origin documents like standards, PNAC/ILAC guidelines etc. (NC) The lab has defined that all the documents are approved by CEO. The lab has not defined the frequency for document review. (NC) | Ex. Audit | - The lab hasn't set up a system to control and manage external documents, like standards and guidelines. They've focused only on internal documents and didn't realize that external documents need to be controlled in the same way to stay compliant and up-to-date. - The root cause of this issue is that the lab has not set a clear rule or process for how often documents should be reviewed. The focus was more on the CEO's approval rather than defining a specific review schedule. | Corrective action for shortcomings on both the NCs has been taken. Lab has revised Master list (QCL-FRM-20.02) and controlled the external origin documents. The lab has revised the procedure (QCL-LSP-07) and defined the frequency for document review. | 17-02-2025 |

| Prepared By: Masab Rasheed AQCM | Evaluated by: Himayat Hussain QCM |
| --- | --- |
| Signature: | Signature: |
| Date: | Date: |