| Sr. No | Non-Conformance | Source | Root Cause | Corrective Action | Target Date |
| --- | --- | --- | --- | --- | --- |
| 01 | The lab has not defined the detail regarding range of lab activities carried out under the accreditation scope ISO/IEC 17025:2017 for testing of fertilizer. Lab needs to define and document all lab activities in its QMS Document (NC) | Ex. Audit | The laboratory considered the scope-extension parameters as proposed activities and planned to incorporate them into the Quality Manual after receiving recommendation/approval from the accreditation body. However, the QMS did not clearly define when proposed scope-extension activities should be incorporated into the relevant documents, who would be responsible for the update, and how the completeness of documentation would be reviewed before submission of the scope-extension application. | The existing QMS documentation was revised to include the requirements for scope-extension activities, including defined responsibility, review of relevant documents, and verification that all proposed parameters are incorporated before assessment. | 10-08-2026 |
| 02 | Laboratory has developed a document vide # QCL-ORG-01, CEO is the top management and QC manager directly report to CEO. The authorization permit of Quality control manager (Mr. Himayat Hussain) Checked the authorization permit needs to revise as per JDs (QCL-JD-02 REV#01 Dated 24-09-24) of the QCM (NC) | Ex. Audit | The authorization permit was not cross-referenced with the latest approved Job Description, resulting in a difference between the defined responsibilities and authorized responsibilities. | The QCM authorization permit was revised and aligned with QCL-JD-02 Rev. 01, and future revisions of JDs will include review of the related authorization permits. | 11-08-2026 |
| 03 | The Lab has defined its quality objectives (QCL/FRM/18.01) for the year 2025 04 quality objectives defined by the Lab with target dates & responsibilities, but the target dates of all objectives are fixed Dec 2025 without any justification and detailed roadmap. (NC) | Ex. Audit | The quality objective planning process did not include a requirement for individual timelines, milestones, and justification of target dates. | QCL/FRM/18.01 was revised to include action plan, milestones, target date justification, and periodic progress review for each quality objective. | 12-08-2026 |
| 04 | The authorization permit of Quality control manager (Mr. Himayat Hussain) Checked the authorization permit needs to revise as per JDs(QCL-JD-02 REV#01 Dated 24-09-24)of the QCM (NC). Mr.Himayat Hussain Quality control Manager vide authorization permit (QCL-FRM-2.01) as per standard clause 6.2.6 clause (a) and authorization related to the interpretation and implitation of the Decision rule/statement of conformity is missing as per clause 6.2.6(b) (NC) | Ex. Audit | The authorization permits (QCL-FRM-2.01) was not completely reviewed and updated after revision of the QCM Job Description (QCL-JD-02 Rev.01). As a result, the required authorization for interpretation and implementation of the decision rule/statement of conformity as per ISO/IEC 17025 clause 6.2.6(b) was inadvertently omitted. | The QCM authorization permit (QCL-FRM-2.01) was revised in accordance with the updated Job Description and ISO/IEC 17025 clause 6.2.6. Authorization for interpretation and implementation of the decision rule/statement of conformity was added, and the revised permit was reviewed and approved. | 20-08-2026 |
| 05 | The lab has developed the Equipment maintenance records sheet (QCL-FRM-4.03) and records related to Equipment maintenance is not maintained (NC) | Ex. Audit | Regular monitoring and verification of equipment maintenance records was not adequately implemented. | Responsibility for maintaining QCL-FRM-4.03 was assigned, and periodic review of equipment maintenance records was included in the laboratory equipment monitoring system. | 21-08-2026 |
| 06 | The lab is using refrigerator for keeping the CRMs and working standards and has also maintained the records of temperature range in between (-5C to +5C) for the but the thermo hygrometer range of calibration/ sensitivity of the thermohygrometer is not intact (NC). | Ex. Audit | The thermo-hygrometer was available and calibrated; however, its calibration range started above 15°C and therefore did not cover the refrigerator operating range of -5°C to +5°C. The suitability of the calibration range for the intended use was not verified before using the instrument for refrigerator temperature monitoring. | A new data logger was purchased because the existing thermo-hygrometer could not be calibrated at the required lower temperature range. The new data logger was calibrated to cover the refrigerator operating range of -5°C to +5°C and was put into use for monitoring the storage conditions of CRMs and working standards. | 22-08-2026 |

| Prepared By: Masab Rasheed AQCM | Reviewed by: Himayat Hussain QCM |
| --- | --- |
| Signature: | Signature: |
| Date: | Date: |