| Sr. No | Non-Conformance | Source | Root Cause | Corrective Action | Target Date |
| --- | --- | --- | --- | --- | --- |
| 01 | Lab has developed a procedure for sample disposal vide Doc# QCL-REG-08 but data missing. (NC) | Internal Audit | The sample disposal procedure was available, but its implementation was ineffective due to lack of unclear responsibility for recording/review, and insufficient monitoring of sample disposal records. | Sample Disposal Register/ Form shall be prepared and controlled. Available sample disposal information shall be updated/reconciled where possible. All relevant staff shall be trained on proper recording requirements. Preventive Action: Monthly review of sample disposal records shall be added to the lab’s routine monitoring checklist and verified during internal audits to ensure continuous compliance. | 17-07-2025 |
| 02 | Blast of KJeldhal’s Assembly Splash head during Nitrogen Testing. | Lab Activity | Lack of standardized verification check list of Kjeldhal’s assembly setup before operation. | Replace damaged splash heads/glassware immediately and the equipment was functionally verified before further use. Prepared pre-use inspection check list and laboratory personnel were instructed on proper equipment handling and operational checks. Additionally, check list display on board for routine verification before testing nitrogen on Kjeldhal’s Apparatus. | 12-05-2026 |

| Prepared by: Masab Rasheed AQCM | Evaluated by: Himayat Hussain QCM |
| --- | --- |
| Signature: | Signature: |
| Date: | Date: |