| Non-Conformance Description |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Blast of KJeldhal’s Assembly Splash head during Nitrogen Testing. |  |  |  |  |  |  |  |  |
| RCA Team Members |  |  | Himayat Hussain (QCM), Masab Rasheed (AQCM) |  |  |  |  |  |
| RCA Team Lead / Head |  |  | Himayat Hussain (QCM) |  |  |  |  |  |
| RCA Start Date |  |  | 11-05-2026 |  | Submission Date |  | 12-05-2026 |  |
| Methodology / Tools Used |  |  |  |  |  |  |  |  |
| Cause & Effect Diagram | □ | 5WHY analysis |  | □ |  | 3W2H / 5W2H |  |  |

Cause Effect Diagram

Documentation

Training

Root Cause

Proper tools

Method

| 3W2H /5W2H | Phenomenon of Problem / Question | Explanation / Answer |
| --- | --- | --- |
| WHY | Why did the splash head blast during nitrogen testing? | Because pressure built up inside the Kjeldhal’s assembly during operation. |
| WHY | Why did pressure build up inside the assembly? | Because the vapor/gas passage was partially obstructed or improperly assembled. |
| WHY | Why was the assembly obstructed or improperly assembled? | Because assembly verification was not performed before starting the test. |
| WHY | Why was assembly verification not performed? | Because there was no structured pre-use operational checklist available. |
| WHY | Why was no structured checklist available? | Because routine procedural controls for equipment setup verification were not formally displayed. |
| How | How did it happen? | Improper assembly or blockage in the Kjeldahl setup caused internal pressure generation, leading to splash head failure during nitrogen testing. |
| How | How to prevent recurrence? | Introduce mandatory assembly verification, pre-use checklist, operator training, and routine inspection before nitrogen testing. awareness/training on equipment handling and inspection requirements. |
| Root Cause: Lack of standardized verification check list of Kjeldhal’s assembly setup before operation. |  |  |

RCA Statement: Lack of standardized verification check list of Kjeldhal’s assembly setup before operation.

Reviewed by: _______________________	                           Evaluated by: ______ ______