| Non-Conformance Description |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| The lab is using refrigerator for keeping the CRMs and working standards and has also maintained the records of temperature range in between (-5C to +5C) for the but the thermo hygrometer range of calibration/ sensitivity of the thermohygrometer is not intact (NC). |  |  |  |  |  |  |  |  |
| RCA Team Members |  |  | Himayat Hussain (QCM), Masab Rasheed (AQCM) |  |  |  |  |  |
| RCA Team Lead / Head |  |  | Himayat Hussain (QCM) |  |  |  |  |  |
| RCA Start Date |  |  | 22-07-2026 |  | Submission Date |  | 22-08-2026 |  |
| Methodology / Tools Used |  |  |  |  |  |  |  |  |
| Cause & Effect Diagram | □ | 5WHY analysis |  | □ |  | 3W2H / 5W2H |  |  |

Cause Effect Diagram

Documentation

Training

Root Cause

Proper tools

Method

| 3W2H /5W2H | Phenomenon of Problem / Question | Explanation / Answer |
| --- | --- | --- |
| WHY | Why was the thermo-hygrometer not suitable for refrigerator monitoring? | Because its calibration range started above 15°C and did not cover the required refrigerator range of -5°C to +5°C. |
| WHY | Why did its calibration range not cover the required lower temperature? | Because the available thermo-hygrometer could not be calibrated at the required lower temperature range. |
| WHY | Why was the same thermo-hygrometer still being used? | Because the limitation of its calibration range was not identified before use. |
| WHY | Why was the limitation not identified? | Because the calibration range was not compared with the intended operating range of the refrigerator. |
| WHY | Why was this comparison not performed? | Because suitability of the measuring equipment for its intended use was not adequately verified during equipment review. |
| How | How did the NC occur? | The available thermo-hygrometer was calibrated only above 15°C, while it was being used to monitor a refrigerator operating between -5°C and +5°C. |
| How | How was the issue addressed/prevented from recurring? | A new data logger was purchased and calibrated for the required lower temperature range, and its suitability was verified before use for monitoring the refrigerator. |
| Root Cause: The thermo-hygrometer was available and calibrated; however, its calibration range started above 15°C and therefore did not cover the refrigerator operating range of -5°C to +5°C. The suitability of the calibration range for the intended use was not verified before using the instrument for refrigerator temperature monitoring. |  |  |

RCA Statement: The thermo-hygrometer was available and calibrated; however, its calibration range started above 15°C and therefore did not cover the refrigerator operating range of -5°C to +5°C. The suitability of the calibration range for the intended use was not verified before using the instrument for refrigerator temperature monitoring.

Evaluated by: ____________                                       Reviewed by: _________________