| NC# | Date: | Non-Conformance | Source | Initiated By |
| --- | --- | --- | --- | --- |
| Final Audit NCs (Team Leader) |  |  |  |  |
| 1 TL | 01-02-2025 | The lab has also developed a format for inspection of purchased products, but has not evaluated as per standard requirements, (NC). Lab has also not defined about the re-evaluation and lab not practicing the reevaluation of vendors (NC). | Ex. Audit | Masab Rasheed |
| 2 TL | 03-02-2025 | The lab has a developed a procedure (QCL/LSP/14) for the non- conforming work. The procedure normally describes about the non-conformities. (NC) The lab has not defined responsibilities and authorities for the management of non-conforming work, (NC). The lab has developed a register for recording the non-conforming work, but lab not checking the significance. (NC) | Ex. Audit | Masab Rasheed |
| 3 TL | 04-02-2025 | The Lab has defined its quality objectives (QCL-FRM-18.01) for the year 2024. 04 quality objectives defined by the Lab with target dates & responsibilities, but the target dates of all objectives are fixed Dec 2024 without any justification and detailed roadmap. (NC) | Ex. Audit | Masab Rasheed |
| 4 TL | 01-02-2025 | The lab has developed a master list (QCL/FRM/20.02) for internal, but lab not controlled the external origin documents like standards, PNAC/ILAC guidelines etc. (NC) The lab has defined that all the documents are approved by CEO. The lab has not defined the frequency for document review. (NC) | Ex. Audit | Masab Rasheed |
| Final Audit NCs (Technical Assessor) |  |  |  |  |
| 01 | 01-02-2025 | The Lab Assistant has not the requisite education, training and competency to conduct the test under scope which is volition of the clause 6.2.2 of the standard ISO/IEC 17025:2017 (NC). Mr. Ahmad Umair (CEO) is supervising the analytical activities in VAN lab as a whole without authorization permit (NC). The lab has developed an annual training Program me (QCL-FRM-2.03) checked the training record of Mr. Masab Rasheed authorized for training of the staff as mentioned in the Authorization permit (QCL/FRM-2.01) but he has no training over MU but working as a trainer over MU (NC). Lab has not defined the criteria of the selection of the trainer (NC). The lab has also developed a training need assessment form (QCL/FRM-2.02) but the criteria of the need assessment not mentioned in procedure as well as in training need assessment farm (QCL-FRM-2.02) (NC). The training record related to Measurement of uncertainty is not available in folder (QCL-FIL-20) (NC) however the records related to trainer evaluation is also missing (NC). | Ex. Audit | Masab Rasheed |
| 02 | 03-02-2025 | The lab is monitoring the temperature and humidity by using thermo hygrometer (VAN/PFTL/EQP-001) in main lab hall while in sample storage room by using the Thermo hygrometer (VAN/PFTL/EQP-002) all are calibrated but the lab has not authorized the person for this activity (NC). The lab has not developed and maintained the log books for analytical weighing balance (QCL-INS-01) and Block heater (QC-INS-023) (NC). | Ex. Audit | Masab Rasheed |
| 03 | 03-02-2025 | All the equipment/apparatus have unique identification number and calibration sticker, wherever required and the equipment logbook analytical balance (QC-INS-001) not maintained (NC). The Calibration plan (QCL-FRM-4.04) for the year 2024 is available in file no (QCL-FIL-22). the lab has developed the Equipment maintenance records sheet (QCL-FRM-4.03) and records related to Equipment maintenance is not maintained (NC). | Ex. Audit | Masab Rasheed |
| 04 | 03-02-2025 | The Lab has developed a procedure related to metrological traceability (QSP-LSP-05) The reference standards used are valid and traceable to their primary standards and stored in refrigerator but monitoring the temperature through un calibrated thermo hygrometer (NC). | Ex. Audit | Masab Rasheed |
| 05 | 04-02-2025 | The lab has developed the simple receiving form (QCL-FRM-12.01) and mentioned the criteria for sample acceptance/rejection like improper sample receiving, costumer/sample information is not appropriate, capability and resources do not meet the requirement, out of the detection limit of the equipment etc. but the acceptance criteria is not maintained in the procedure, lab needs to revise the procedure (QCL-LSP-22) (NC). | Ex. Audit | Masab Rasheed |
| 06 | 07-02-2025 | The lab has not developed a procedure for validation and verification of methods (NC). The raw data regarding the method validation is not maintained (NC). | Ex. Audit | Masab Rasheed |
| 07 | 08-02-2025 | The VAN Laboratory has not developed a lab procedure for control of technical record (QCL-LSP-20) (NC). | Ex. Audit | Masab Rasheed |
| 08 | 10-02-2025 | The lab analyst involved in reporting of Uncertainty of measurements have not sufficient knowledge over MU calculation lab staff involved in the calculation of MU needs training (NC). The raw data related to uncertainty measurements Nitrogen Phosphorus is not properly available (NC) | Ex. Audit | Masab Rasheed |
| 09 | 10-02-2025 | The lab has developed the procedure for reporting of results (QCL-LSP-19) however lab needs to revise the procedure as per Clause 7.8.1 & 7.8.2 of the standard ISO/IEC 17025:2017 (NC). | Ex. Audit | Masab Rasheed |
| 10 | 11-02-2025 | Laboratory has not applying the decision rule and also not mentioned lab procedure of reporting of results (NC). Lab has not applying the decision rule for the statement of conformity (NC). The VAN laboratory is providing interpretation like ok but the criteria regarding ok and not ok not mentioned in procedure (QCL-LSP-19) (NC). The lab has a procedure for reporting of results but not mentioned about amended/duplicate report issuance not mentioned in procedure (NC). | Ex. Audit | Masab Rasheed |
| Internal Audit NC |  |  |  |  |
| 01 | 17-06-2025 | Lab has developed a procedure for sample disposal vide Doc# QCL-REG-08 but data missing.NC | Internal Audit | Masab Rasheed |
| Lab Activity NC |  |  |  |  |
| 02 | 11-05-2026 | Blast of Kjeldhal’s Assembly Splash head during Nitrogen Testing. | Lab Activity | Masab Rasheed |
| Surveillance and Scope Extension NCs (Team Leader) |  |  |  |  |
| 01 TL | 15-07-2026 | The lab has not defined the detail regarding range of lab activities carried out under the accreditation scope ISO/IEC 17025:2017 for testing of fertilizer. Lab needs to define and document all lab activities in its QMS Document (NC) | Ex. Audit | Masab Rasheed |
| 02 TL | 16-07-2026 | Laboratory has developed a document vide # QCL-ORG-01, CEO is the top management and QC manager directly report to CEO. The authorization permit of Quality control manager (Mr. Himayat Hussain) Checked the authorization permit needs to revise as per JDs (QCL-JD-02 REV#01 Dated 24-09-24) of the QCM (NC) | Ex. Audit | Masab Rasheed |
| 03 TL | 17-07-2026 | The Lab has defined its quality objectives (QCL/FRM/18.01) for the year 2025 04 quality objectives defined by the Lab with target dates & responsibilities, but the target dates of all objectives are fixed Dec 2025 without any justification and detailed roadmap. (NC) | Ex. Audit | Masab Rasheed |
| Surveillance and Scope Extension NCs (Technical Assessor) |  |  |  |  |
| 01 TA | 20-07-2026 | The authorization permits of Quality control manager (Mr. Himayat Hussain) Checked the authorization permit needs to revise as per JDs (QCL-JD-02 REV#01 Dated 24-09-24) of the QCM (NC). Mr. Himayat Hussain Quality control Manager vide authorization permit (QCL-FRM-2.01) as per standard clause 6.2.6 clause (a) and authorization related to the interpretation and implitation of the Decision rule/statement of conformity is missing as per clause 6.2.6(b) (NC) | Ex. Audit | Masab Rasheed |
| 02 TA | 21-07-2026 | The lab has developed the Equipment maintenance records sheet (QCL-FRM-4.03) and records related to Equipment maintenance is not maintained (NC) | Ex. Audit | Masab Rasheed |
| 03 TA | 22-07-2026 | The lab is using refrigerator for keeping the CRMs and working standards and has also maintained the records of temperature range in between (-5C to +5C) for the but the thermo hygrometer range of calibration/ sensitivity of the thermohygrometer is not intact (NC). | Ex. Audit | Masab Rasheed |

Prepared By: ________________								Updated on: 22-07-2026