| NC# | Date: | Non-Conformance | Source | Initiated By |
| --- | --- | --- | --- | --- |
| Surveillance and Scope Extension NCs (Team Leader) |  |  |  |  |
| 01 TL | 15-07-2026 | The lab has not defined the detail regarding range of lab activities carried out under the accreditation scope ISO/IEC 17025:2017 for testing of fertilizer. Lab needs to define and document all lab activities in its QMS Document (NC) | Ex. Audit | Masab Rasheed |
| 02 TL | 16-07-2026 | Laboratory has developed a document vide # QCL-ORG-01, CEO is the top management and QC manager directly report to CEO. The authorization permit of Quality control manager (Mr. Himayat Hussain) Checked the authorization permit needs to revise as per JDs (QCL-JD-02 REV#01 Dated 24-09-24) of the QCM (NC) | Ex. Audit | Masab Rasheed |
| 03 TL | 17-07-2026 | The Lab has defined its quality objectives (QCL/FRM/18.01) for the year 2025 04 quality objectives defined by the Lab with target dates & responsibilities, but the target dates of all objectives are fixed Dec 2025 without any justification and detailed roadmap. (NC) | Ex. Audit | Masab Rasheed |
| Surveillance and Scope Extension NCs (Technical Assessor) |  |  |  |  |
| 01 TA | 20-07-2026 | The authorization permits of Quality control manager (Mr. Himayat Hussain) Checked the authorization permit needs to revise as per JDs (QCL-JD-02 REV#01 Dated 24-09-24) of the QCM (NC). Mr. Himayat Hussain Quality control Manager vide authorization permit (QCL-FRM-2.01) as per standard clause 6.2.6 clause (a) and authorization related to the interpretation and implitation of the Decision rule/statement of conformity is missing as per clause 6.2.6(b) (NC) | Ex. Audit | Masab Rasheed |
| 02 TA | 21-07-2026 | The lab has developed the Equipment maintenance records sheet (QCL-FRM-4.03) and records related to Equipment maintenance is not maintained (NC) | Ex. Audit | Masab Rasheed |
| 03 TA | 22-07-2026 | The lab is using refrigerator for keeping the CRMs and working standards and has also maintained the records of temperature range in between (-5C to +5C) for the but the thermo hygrometer range of calibration/ sensitivity of the thermohygrometer is not intact (NC). | Ex. Audit | Masab Rasheed |

Prepared by: ________________								Updated on: 22-07-2026