# Standard Operating Procedure for

Test Methods and Method Validation

ISO/IEC 17025:2017 – Clause 7.2

| Prepared By: Masab Rasheed | Reviewed By: Himayat Hussain | Approved By: Ahmed Umair |
| --- | --- | --- |
| AQCM | QCM | CEO |
| Date: | Date: | Date: |

#### Copyright

The copyright of this procedure, together with all confidential information contained herein is the sole property of VAN QC LAB.

It may be copied in full or in parts only by the Management / personnel and only for the purpose of Company related activities. Disclosure and copying of any information contained within this procedure to any person (s) outside the employee of company without written permission/authorization of the Chief Executive Officer is strictly prohibited.

Amendment Sheet

| Date | Revision # | Amendment Detail | Section | DCR | Amendment Approved by |
| --- | --- | --- | --- | --- | --- |

- Purpose

This document establishes guidelines for method selection and the procedures for verification of standard / validation of non-standard method performance. Standard Testing Methods that do not need validation, but for the confidence of technical staff and commitment to quality, precision and accuracy of results laboratories verify adopted standard testing methods prior to use.

- Application

This procedure covers verification of standard / validation of non-standard testing methods introduced into the Laboratories.

- Terms and Definitions

  - Accuracy: Accuracy is the nearness of a result or the mean of a set of measurements to the true value.

  - Analytical batch: An analytical batch consists of samples which are analyzed together with the same method sequence and the same lots of reagents and with the manipulations common to each sample within the same time or in continuous sequential time periods.

  - Detection limit: A detection limit is the lowest amount of analyte in a sample which can be detected but not necessarily quantitated as an exact value. It is often called the limit of detection (LOD) which is the lowest concentration level that can be determined statistically different from a blank at a specified level of confidence. It is determined from the analysis of sample blanks. Method detection limit (MDL) is the minimum concentration of a substance that can be measured and reported with 99% confidence that the analyte concentration is greater than zero. It is determined from analysis of a sample in each matrix containing the analyte.

  - Limit of Quantitation: This is the level above which quantitative results may be determined with acceptable accuracy and precision.

  - Linearity: Linearity is the ability of the method to elicit results that are directly proportional to analyte concentration within a given range.

  - Non-standard method: This refers to a method that is not taken from authoritative and validated sources. This includes methods from scientific journals and unpublished laboratory-developed methods.

  - Precision: Precision is the agreement between a set of replicate measurements without assumption of knowledge of the true value. Repeatability expresses the precision under the same operating conditions over a short period of time. Intermediate precision expresses within-laboratory variations, such as different days, different analysts, and different equipment. Reproducibility expresses the precision between laboratories.

  - Range: A range is the interval between the upper and lower concentration of analyte in sample for which it has been demonstrated that the analytical procedure has an acceptable level of accuracy, precision, and linearity.

  - Ruggedness or robustness: Ruggedness is a measure of an analytical procedure’s capacity to remain unaffected by small, but deliberate variations in method parameters and provides an indication of its reliability during normal usage.

  - Specificity: Specificity is the ability to assess unequivocally the analyte in the presence of components which may be expected to be present.

  - Lab Standard Procedure (LSP): This is a method that is traceable to a recognized, validated method.

  - Method Validation: A method validation is the process of establishing the performance characteristics and limitations of a method and the identification of the influences which may change these characteristics and to what extent.

  - Verification: Verification is the confirmation by examination and provision of objective evidence that specified requirements have been fulfilled.

- Responsibility

  - The AQCM have the responsibility to ensure implementation of method verification procedures.

  - QCM has the responsibility to implement method verification procedures in Lab.

- Procedure

  - Method Selection

An integral part of the laboratory quality system is the use of standard testing methods. Non-standard methods could also be used after validation.

  - Standard Method Verification

    - Standard methods need verification to ensure that the laboratory can perform the analysis. The laboratory’s quality control program as described in Quality System Manual. Verification of an analytical method is the demonstration that a laboratory is capable of replicating with an acceptable level of performance of a standard method. Verification under conditions of use is demonstrated by meeting system suitability specifications established for the method, as well as a demonstration of accuracy and precision or other method parameters for the type of method. Method performance is accomplished by using performance characteristics such as:

      - Precision based on analysis of multiplicities

      - Accuracy based on analysis of samples of known concentrations (Certified Solutions/Reference Material)

      - Monitoring quality control samples, usually using Range Control Charts.

    - To carry out method verification a team of technical staff members under the supervision of AQCM is formed. The team directly reports to QCM as a team leader.

    - Team selects sample(s) of known value for method verification, and checks that the laboratory has:

      - Suitable/required equipment.

      - Chemicals/standards/solutions.

      - Appropriate environment.

      - Competent personnel for performing the standard method.

    - After identifying and acquiring resources required to execute the assignment the Team carries out method verification by following the standard testing method being verified.

    - Details of analysis including accuracy, precision, any special instructions for the use of standard testing method and a statement of weather the method has been adopted for use or will require certain modifications are documented on Standard Test Method Verification Form.

    - Accuracy is determined by analyzing % Recovery as:

% Recovery = (X / K) * 100

Where X: Measured Value

K: Known value

    - Precision is expressed as Relative Standard Deviation (RSD) or Relative Percent Difference (RPD) of multiple samples.

RSD is calculated from standard deviation and mean recovery, when the standard deviation is derived from multiple recovery results as follows:

RSD = CV = (σ / x) * 100

Where,

RSD = Relative Standard Deviation

CV = Coefficient of variation

σ = SD (Standard Deviation) = { Σ (xi – x)2 / (n – 1) }1/2

x  = Mean

xi  = Measured Value (x1, x2, x3, …, xn)

n  = No of measurements

RPD is calculated when only two sample results are available as follows:

RPD = { ( | R1 - R2 | ) / x } * 100

Where:

| R1 - R2 | = absolute difference between the determinations

    - In cases where the SOP /WI is modified from the standard or existing testing procedure and protocol, it is demonstrated by the Team that the modifications do not adversely affect the precision and accuracy of the data obtained.

    - To implement the modification, the standard or existing method is first performed. Each major modification is then validated against the original method.

    - Additional step to verify performance specification is ‘Mean Test’ for significance of difference between the two samples means to determine degree of accuracy. The sample mean of readings taken by existing method must be close to readings taken by modified method while reading taken by modified method must be close and in an acceptable range of the true/known value of the sample.

    - The statement on verification of standard testing methods also includes a note about acceptability and details of modifications made to the standard testing method(s) if any.

In addition to determining the accuracy and precision the factors like number of operators, number of equipment and reproducibility with respect to time is taken into consideration for determining verification of the method keeping in view the tolerance limit of the product to be tested.

  - Method Validation

In addition to the parameters defined in 5.2, following parameters are required:

    - Linearity

### The linearity of an analytical procedure is its ability (within a given range) to obtain test results which are directly proportional to the concentration (amount) of analyte in the sample

For the establishment of linearity, minimum 5 concentrations are recommended (from LOQ till the range last or within specified range). The graph of concentration of analyte vs. response must be plotted. The independent variable (concentration) must be plotted against X-axis and dependent variable (i.e., response) must be plotted against Y-axis.

    - LOD / LOQ

The limit of detection LOD (or detection limit, DL) is the lowest possible concentration at which the method can detect (but not quantify!) the analyte within the matrix with certain degree of confidence.

LOD is calculated in the following way:

LOD = 3.3 * σ / S

Where,

σ = SD of the obtained results

S = Slope of the calibration curve.

The limit of quantitation LOQ (or sometimes also referred as quantification limit, QL) is the lowest possible concentration of the analyte that can be quantified by the method in a reliable way.

LOQ is calculated as follows:

LOQ = 10 * σ / S

Where,

σ = SD of the obtained results

S = Slope of the calibration curve.

    - Ruggedness and Robustness

The robustness/ruggedness of an analytical procedure is a measure of its capacity to remain unaffected by small, but deliberate variations in method parameters and provides an indication of its reliability during normal usage. The guidelines below shall be followed for calculations

https://fabi.research.vub.be/sites/default/files/atoms/files/robustness_tests.pdf

    - PT Participation

Participation in the same or similar PT program and qualifying with a satisfactory Z-score is mandatory.